## Announcement Differential pressure transmitter P26

Configurable pressure transmitter with

automatic zero-point correction

- Measurement ranges from 0..10 Pa to 100 kPa

- For measuring positive or negative pressures, even asymmetrically

- Scalable measurement ranges and units

- Measurement ranges / outputs can be configured

- Automatic zero-point calibration prevents zero-point drift

- Built-in valve provides a high level of overload protection

- Inductive measurement cell for highly accurate measurement and temperature stability

- Can be calibrated (ex works and on-site)

- Also for volume / mass flow and flow rate measurements

with immediate display of the values

- Multilingual operating menu

- With 2 alarm contacts

- With multicolour LCD

- USB port

- Top hat rail assembly and wall mounting, IP 65

- CE and CSA approved

Technical data

Measurement

ranges:	10/ 50/ 100/ 250/ 500 Pa, 1 / 2.5 / 5/ 10 / 20 / 50 / 100 kPa (others available upon request) freely scalable from 10..100% within a measurement range

Analog outputs: 	0..10 V, 0..20 mA, 
4..20 mA, ± 5 V

Power supply:	24 VAC, 24 VDC, 
24 VAC, 230/115 VAC with galvanic separation

LCD: 	Multicolour + keyboard

Contact points: 	2 switching relays
max. 230 VAC, 6 A

Margin of error: 	0.2% of the scaled range, min. 0.3 Pa Standard: 0.5% of scaled range

Temperature
coefficient margin:	0.03 % of max. value/K

Temperature 
coefficient

zero point:	± 0.0 %

Max. system pressure

Overload capacity: 	600 kPa for measurement ranges ≥2.5 kPa

200x for measurement ranges <2.5 kPa

Medium: 	Air, all non-aggressive gases

Sensor 
response time: 	25 ms

Time constant: 	20 ms.. 60 s (adjustable)

Operating

temperature:	10°C .. 50°C,

Storage

temperature: 	-10°C .. 70 °C

Power
consumption: 	approx. 6 VA

Weight: 	approx. 0.75 kg

Cable glands: 	3 x M 16

Pressure ports: 	Nominal width of 6 mm for hose, others available upon request

Protection class: 	IP 65, with USB IP 40

Housing/cover: 	ABS

Dimensions: 	164 mm x 102.5 mm x 76 mm

Wall mounted

housing: 	Top hat rail assembly and wall mounting possible

Certificates: 	CE, CSA

Data interface: 	USB, data cable supplied

External zero-point calibration

Optional: 	DAkkS calibration certificate,

ISO factory calibration certificate

Designation:	Differential pressure transmitter

Type: 	P26

Manufacturer: 	halstrup-walcher GmbH